|  | - University of Liverpool Radiation Protection Office THE USE OF IONISING RADIATION IN UNDERGRADUATE TEACHING |
| --- | --- |

INTRODUCTION

The following guidance notes are for departments that use ionising radiation in undergraduate teaching.

The University’s Radiation Protection Adviser (RPA) recommends that the radiation protection standards laid down for members of the public should be adhered to in general, with a small number of exceptions permitted in the case of 3rd year students whose work is directly supervised by a member of staff.

The whole body (effective) dose limit for members of the public and undergraduates is 1 mSv y-1 (above background radiation). In addition, it is recommended that the equivalent dose to the lens of the eye from each lab teaching course should not exceed 5 mSv and for extremities should be less than 10 mSv.

These limits are quite low and may be compared with the background dose rate of approximately 2.7 mSv y-1 and the IRR17 recommended radiation worker annual dose limit of 20 mSv y-1.

In following this recommendation, the University should not need to register undergraduates as radiation workers or to provide radiation dosimeters. Nevertheless it may choose to do so.

To achieve the standards laid down, all experiments and demonstrations involving the use of ionising radiation should be carefully prepared with a view to minimising radiation exposure at all times (ALARP).  Adequate supervision of undergraduates must be provided and instruction in correct handling techniques with an explanation of the hazards involved should precede the actual experiment.

A summary of the main conditions to ensure compliance is given below:

X-RAY APPARATUS

Should be completely enclosed (including the main beam) in protective shielding such that the Instantaneous Dose Rate dose not exceed  5 μSv h-1 at a distance of 5cm from the shielding at the maximum operating conditions of the set.  Interlocks should be provided to shut off the useful beam if the shielding is moved.

DISCHARGE TUBES

These are potential sources of high radiation dose rates and should be checked by the University RPA under the actual working conditions before use in demonstration lectures, etc.

CLOSED (SEALED) RADIOACTIVE SOURCES

For sources causing external radiation, the activity is restricted according to the dose rate they cause when fully exposed.  Table I below gives the maximum permitted dose rate at 0.1m and the typical maximum activity for different types of source.

| Source | Maximum Dose Rate μSv h-1 | Typical Maximum Source Activity |
| --- | --- | --- |
| Beta-particle Gamma-ray Neutron | 50 10 10 | 50 kBq 500 kBq 5000 n s-1 |

Table I	Recommended Maximum Dose Rate at 0.1m

Sources that do not satisfy these requirements should be shielded so that the dose rate does not exceed that given above for X-ray apparatus.  The RPA should monitor them before they are used in teaching.  The shielding should carry a warning notice to the effect that a radioactive source is present and the radioisotope and activity involved should be clearly marked.

All sources used in a teaching laboratory must be returned to secure storage at the end of the teaching period and a check made that none are missing.

UNSEALED RADIOACTIVE MATERIAL (OPEN SOURCE)

The laboratory instruction sheet for each experiment should include a description of the potential hazards and the precautions to be taken, and should form the written system of work for that experiment.

The attention of the student should be drawn to the necessity of avoiding contamination of the laboratory or themselves and the usual precautions for the use of open source material should be employed, i.e. Benchkote, spill trays, disposable gloves, lab coats, etc. An abbreviated copy of the General Local Rules may be useful in the Departmental Student Handbook

To reduce to an absolute minimum the risk from inadvertent inhalation or ingestion of unsealed radioactive material, it is recommended that the maximum activity in use in an experiment should not exceed one-tenth of the Annual Limit of Intake (ALI) for a radiation worker of the radioisotope involved.  Table II lists the value of this quantity for the radioisotopes most commonly used in the University of Liverpool.

Students should never be permitted to dispense radioactivity from stock solutions.

| Radionuclide | H3 | C14 | S35 | P32 | Ca45 | I125 |
| --- | --- | --- | --- | --- | --- | --- |
| One-tenth ALI (MBq) | 100 | 4 | 3 | 0.5 | 0.5 | 0.1 |

TABLE II        Maximum Permissable Activity per Experiment

Most undergraduate experiments can be carried out with activities much smaller than those listed above.

Where 3rd year students are engaged in project work under the close supervision of a member of staff, they may handle larger activities of unsealed radioactive material than those listed in Table II, provided the RPA is consulted beforehand on the arrangements for the experiment in question.

Great care must be exercised when handling environmental samples containing natural uranium and thorium compounds. Table III gives the maximum quantity of these materials which may be used in teaching experiments.  The value quoted for uranium takes account of both chemical toxicity and radiation protection requirements.

| Element | Activity (Bq) | Mass (Milligram) |
| --- | --- | --- |
| Thorium | 5 | 1 |
| Uranium | 50 | 5 |

TABLE III        Maximum Activity of Thorium or Uranium per Experiment