|  | # University of Liverpool RISK ASSESSMENT FOR THE USE OF NON-IONISING RADIATION (Ultra Violet) |
| --- | --- |

| Ref No of Assessment | RA_UV_CAMAG_Hand-Held Lamp |
| --- | --- |
| Department |  |
| Location & Room No of Laboratory |  |
| Name & Status of Assessor (see Note 1) | Pete Cole – RPA |
| Name of Departmental Safety Officer |  |
| Make and Model of Equipment | Hand-held CAMAG UV lamp – dual wavelength 8W tubes |
| Serial Number of Equipment | 12345 |

| Identify the Hazard(s) | (delete as applicable) |  |  |
| --- | --- | --- | --- |
| Ultra Violet Radiation | UVA (350nm) | UVB (280-315nm) | UVC (254nm) |

Describe the process or work procedure

Thin Layer Chromatography (TLC)

Who might be affected by this work? (see Note 2)

Research staff and students

| What is the nature of the risk? |  | Short term/long term? | Mild/severe? |
| --- | --- | --- | --- |
| ## Eye | Photo-keratitis / Photo-conjunctivitis | S | S |
|  | Photochemical cataract | L | S |
|  | Other – e.g. possible photochemical retinal challenge | L | M |
| ## Skin | Erythema | S | S |
|  | Ageing | L | M |
|  | Photochemical skin reactions | L | M |
|  | Cancer (including SSC, BSC and MM) | L | S |

| Control Measures to minimise the UV risk | In Use? (Y/N) | Comments / Actions Required |
| --- | --- | --- |
| ## Interlocking: e.g. doors, enclosures etc. | N | Hand-held lamp. |
| Warning Signs and Labels: | Y | “UV” warning label displayed on unit which includes the British Standard “!” hazard warning symbol. |
| Shielding and PPE (to BS169): | Y | All persons within ½ m must make appropriate use of a suitable UV safety full face visor (marked as BS EN 169:2002 compliant) and UV-protective clothing. |
| Training (documented?): | Y | Operators trained in-house for use of this type of device. All operators must read the UV Local Rules. |
| ## Controlled Access: | Y | Swipe-card entry into suite of labs. |
| ## Local Rules (plus Contingency Plan): | Y | UV Local Rules available in lab. |

| Other Hazards | Comments | Control Measures | Actions Required |
| --- | --- | --- | --- |
| Chemical |  |  |  |
| Electrical | Electrical malfunction (see Fire below) | Unit PAT tested annually and marked as such. | Continue annual PAT |
| Biological | NA | NA | NA |
| Fire | Fire caused by electrical malfunction. | CO2 fire extinguishers located within easy access of lab. Tested annually. | Continue annual tests on extinguishers |
| Other | NA | NA | NA |

Assessor's Additional Comments:

Spectral integration for UVC line done over 249 – 259nm, and for UVA line done over 310 – 400nm. Exposure Limit Values = 10,000 J/m2 (total) and 30J/m2 (Effective) Relative spectral effectiveness – S(lambda) from “ICNIRP Guidelines of Exposure to UV Radiation of Wavelengths between 180nm and 400nm (Incoherent Optical Radiation)”, Health Physics, 87(2), 2004, pp 171-186. Measurements Solatell Sola-Scope 2000 (s/n:00096). Solatell Sola-Sensor 2000 (s/n:00120) At a distance of 10cm (potential eye-face position) with no PPE or shielding: UVC (254nm) Line - Effective Irradiance = 0.22 W/m2 Exposure time to Eff. Radiant Exposure of 30J/m2 = 2.3 minutes UVA (350nm) Line - Effective Irradiance = 0.0003 W/m2 There is no significant risk to the eyes or skin from the 350nm mode of operation for the short periods of time that it may be used each day. However, with the 254nm (UVC) mode, both skin and eyes can be damage within a few minutes. Equipment must only be used by suitably trained operators. All operators must read the UV Local Rules available in the lab. The lamp must never be pointed at a person. Assuming that the lamp is used for up to one hour in any one day by the same person, then any user whose hands or eyes are within approximately 50cm of the lamp must make appropriate use of suitable and sufficient PPE when the lamp is switched on. This includes a full face visor and a combination of gloves and lab-coat to cover all areas of bare skin. Full face visor – the filter must be compliant with BS EN 169:2002 (or its US equivalent) and be marked as such. It is suggested that the filter shade should have a scale number of at least ‘2-1,2’ which provides an Optical Density (OD) of approximately OD5.5 up to 313nm and OD0.3 (50% transmission) up to 365nm.

| Assessor's Signature | Printed Name | Date |
| --- | --- | --- |
|  | Prof. Pete Cole | 05/10/18 |

| Reviewer's Signature | Printed Name | Date |
| --- | --- | --- |

Further Notes: To be reviewed in 2021 or sooner if there is ever a significant accident, incident or untoward occurrence.