|  | # University of Liverpool RISK ASSESSMENT FOR THE USE OF NON-IONISING RADIATION (Ultra Violet) |
| --- | --- |

| Ref No of Assessment | RA_UV_Spectroline_Transilluminator |
| --- | --- |
| Department |  |
| Location & Room No of Laboratory |  |
| Name & Status of Assessor (see Note 1) | Pete Cole – RPA |
| Name of Departmental Safety Officer |  |
| Make and Model of Equipment | Spectroline 1C-312A/1/transilluminator x 76016 |
| Serial Number of Equipment | 262389 |

| Identify the Hazard(s) | (delete as applicable) |  |  |
| --- | --- | --- | --- |
| Ultra Violet Radiation | UVA (315-400nm) | UVB (280-315nm) | UVC (100-280nm) |

Describe the process or work procedure

Visualization of protein, DNA, RNA and their precursors in electrophoresis gel plates.

Who might be affected by this work? (see Note 2)

Research staff and students

| What is the nature of the risk? |  | Short term/long term? | Mild/severe? |
| --- | --- | --- | --- |
| ## Eye | Photo-keratitis / Photo-conjunctivitis | S | S |
|  | Photochemical cataract | L | S |
|  | Other – e.g. possible photochemical retinal challenge | L | M |
| ## Skin | Erythema | S | S |
|  | Ageing | L | M |
|  | Photochemical skin reactions | L | M |
|  | Cancer (including SSC, BSC and MM) | L | S |

| Control Measures to minimise the UV risk | In Use? (Y/N) | Comments / Actions Required |
| --- | --- | --- |
| ## Interlocking: e.g. doors, enclosures etc. | N | No micro-switch interlocks installed. Unit fitted with Perspex lid shielding. |
| Warning Signs and Labels: | Y | Warning labels displayed which incorporate the British Standard “!” hazard warning symbol. |
| Shielding and PPE (to BS169/170): | Y | Unit fitted with Perspex lid shielding. Must be NO direct viewing with lid open unless a suitable UV safety full face visor (marked as BS EN 169:2002 compliant) and UV-protective clothing are used. |
| Training (documented?): | Y | Operators trained in-house for use of this type of device. All operators must read the UV Local Rules. |
| ## Controlled Access: | Y | Door to lab not routinely locked but swipe-card entry onto to floor. |
| ## Local Rules (plus Contingency Plan): | Y | UV Local Rules available in lab. |

| Other Hazards | Comments | Control Measures | Actions Required |
| --- | --- | --- | --- |
| Chemical | Use of Ethidium Bromide – potential mutagen. Highly toxic by inhalation. Irritant for eyes, skin and upper respiratory tract. | See separate CoSHH assessment. Suitable PPE used i.e. lab coats, nitrile gloves, safety glasses or visor, etc. | Follow relevant sections of lab safety procedures. |
| Electrical | Electrical malfunction (see Fire below) | Unit PAT tested annually and marked as such. |  |
| Biological |  |  |  |
| Fire | Fire caused by electrical malfunction. | CO2 fire extinguishers located within easy access of lab. Tested annually. |  |
| Other |  |  |  |

Assessor's Additional Comments:

Spectroline 1C-312A Transilluminator 0 0.00001 0.00002 0.00003 0.00004 0.00005 0.00006 230 280 330 380 430 Wavelength (nm) Spectral Intensity (W/cm2/nm) Spectral integration done over 272 – 400nm. ELV = 10,000 J/m2 (total) and 30J/m2 (Effective) Relative spectral effectiveness – S(lambda) from “ICNIRP Guidelines of Exposure to UV Radiation of Wavelengths between 180nm and 400nm (Incoherent Optical Radiation)”, Health Physics, 87(2), 2004, pp 171-186. Measurements Solatell Sola-Scope 2000 (s/n:00096) Solatell Sola-Sensor 2000 (s/n:00120) At a distance of 13cm (typical eye-face position) with shielding lid open (i.e. no shielding): Total Irradiance = 14.1 W/m2 Effective Irradiance = 1.8 W/m2 Exposure time to Eff. Radiant Exposure of 30J/m2 = 16 s With shielding lid closed (shielded): Effective Irradiance = 0.0002 W/m2 With the lid shield closed there is no risk to the eyes or skin. However, with the lid open (i.e. unshielded UV source) both skin and eyes can be damage within a few seconds. Whenever reasonably practicable the unit must only be switched on when the lid is fully closed. If the unit is ever to be operated with the lid open then operators MUST make use of suitable and sufficient PPE. This includes a full face visor and a combination of gloves and lab-coat to cover all areas of bare skin. Full face visor – the filter must be compliant with BS EN 169:2002 (or its US equivalent) and be marked as such. It is suggested that the filter shade should have a scale number of at least ‘2-4’ which provides an Optical Density (OD) of approximately OD5.5 up to 313nm and OD2 up to 365nm.

| Assessor's Signature | Printed Name | Date |
| --- | --- | --- |
|  | Dr. Pete Cole | 05/02/07 |

| Reviewer's Signature | Printed Name | Date |
| --- | --- | --- |

Further Notes: To be reviewed in 2021 or sooner if there is ever a significant accident, incident or untoward occurrence.